SEM Hitachi S4700 User Manual

General checks

- Write “Date”, “User” and “Start time” on a user log

- Check vacuum (IP1: 1.0 x 10-8 Pa, IP2: 1.0 x 10-7 Pa, IP3: 1.0 x 10-6 Pa, S.C.(Pe): L x 10-3 Pa, S.C.(Pi): L x 10-1 Pa, S.E.C.(Pi): 1.0 Pa)

- Turn on the PC display

- Double click the SEM observation software

- If “please flashing” come out in the scanning image window, flash the electron gun

Sample insertion

- Open nitrogen valves (the nitrogen bomb is located behind the SEM)

- Fill up balloon with nitrogen

- Push the “Air” button on front panel of the SEM to vent the specimen exchanging chamber (SEC)

- Close nitrogen valves

- Pull the door open (DO NOT OPEN the MV (manual valve of specimen chamber) by grabbing the SEC, not the rod (do not touch inside of SEC and USE GLOVE)

- Push the specimen exchange rod slightly to unlock it (use glove)

- Unscrew the sample holder onto the end of the rod (use glove)

- Put sample on the sample holder (use glove)

- Check the sample height level (use glove)

- Screw the sample holder onto the end of the rod (use glove)

- Pull the rod back to the locked position (use glove)

- Bring the SEC unit to SEC-SC interface (use glove)

- Press the “EVAC” button to evacuate the SEC with HOLDING the SEC unit

- Wait until vacuum becoming lower than 2.0 Pa

- Open the MV and insert the sample holder by pushing the specimen exchange rod until it stopped completely

- Unscrew the sample holder

- Make sure the sample holder is completely released from the rod and mount on the receiver

- Pull the rod back to the locked position (carefully and slowly)

- Close the MV

Observation

- Pull up the lever for gun valve on front panel to be “Auto”

- Turn on HV

- Optimize focus and stigma (~50K in magnification is good for optimization)

- Click alignment which is a fourth icon from left on top in the window

- Optimize beam alignments

- Observe your sample

- When you capture SEM images, change scan speed to be “S2” and push “CAP” button

Sample withdrawal

- Turn off HV

- Push down the lever for gun valve on front panel to be “Close”

- Move the stage position to the sample exchange position (X, Y stage: 12.5, X-Y Rotation: 0 deg, Tilt: 0 deg, Z position: 12.0)

- Open the MV (use glove)

- Insert the rod until it stopped completely

- Screw the rod into the sample holder

- Make sure it is completely screwed

- Pull the rod back to the locked position (carefully and slowly)

- Close the MV

- Open nitrogen valves

- Fill up the balloon with nitrogen

- Push the “Air” button to vent the SEC

- Close nitrogen valves

- Unlock the specimen exchange rod (use glove)

- Unscrew and pick up sample holder from the rod (use glove)

- Pick up your sample from the sample holder (use glove)

Clean up

- Remove a top part of the sample holder (carbon tape would be bad for vacuum condition in specimen exchange chamber)

- Screw the sample holder onto the end of the rod (use glove)

- Pull the rod back to the locked position (use glove)

- Bring the SEC unit to SEC-SC interface (use glove)

- Press the “EVAC” button to evacuate the SEC with HOLDING the SEC unit

- Close the SEM observation software

- Turn off the PC display (not PC power, but only display)

- Write “End time”, “Vext” and “OK (if you did not see any problem in your SEM observation)” on the user log

- Clean up everything (tweezers, carbon tape, scissors and so on) on the desk