Raman Microscope inVia Reflex User Manual

Stand by

- Write the user record (Date, Start time, Affiliation, User name, Sample, Laser and Grating)

- Turn on main power of the main body

- Turn on the laser power (532 nm) switch not the key switch. Wait to be lightening the “Stable Emission” lamp (it takes 2-3 min), then rotate the key switch to be “ON”

- Go to Raman PC

- Double click “Wire 3.2”

- Choose “Reference un-referenced motors only” and click “OK” (if the window come out)

- Wait for 25 min to cool down the photo detector

- Remove the plastic cover on the Raman microscope

Calibration by using silicon (Si) sample

- Put Si sample (in the box next to the microscope) on the optical microscope slide

- Chose the x5 objective lens by rotating the revolver (*do not grab the lens tube)

- Move the sample under the objective lens by sliding the stage

- Confirm or turn on the camera and light bulb by clicking the camera icon (View the sample under Video And Eyepieces) and the light bulb icon (Sample Illumination), respectively

- Roughly focus on the sample by rotating the Z position dial

- Choose the x20 objective lens by rotating the revolver

- Accurately focus on the sample by using the ball type mouse (you can use the laser spot (laser power<5x10-5))

- Click the “Measurement  New  Spectral acquisition”

- Select or enter the “Static” (grating scan type), “Standard” (confocality), the 520 as center (spectrum range), 532 nm (laser name), and 1800 l/mm [vis] (grating name)

- Go to the “Acquisition” tab

- Set the 1.00 s (Exposure time), 1 (accumulation), and 1% (laser power).

- Check the box for the “Close laser shutter on completion”

- Click “OK”

- Close the curtain and turn off the room light

- Click the “Run icon”

- Check the peak position (~520.5 cm-1) of the Si

[If the Si peak was far from 520.5 cm-1]

- Go to “Tool” tab  Calibration  Offset  Enter the offset value (i.e. if measured peak was 517.5 cm-1, enter the -3 cm-1 then OK)

- Click the “Run icon” again to make sure the calibration was done

- File  Save file as  name your data  save as wxd. file

- Take out the Si sample

Measurement

- Put your sample (in the box next to the microscope) on the optical microscope slide

- Chose the x5 objective lens by rotating the revolver (*do not grab the lens tube)

- Move the sample under the objective lens by sliding the stage

- Confirm or turn on the camera and light bulb by clicking the camera icon (View the sample under Video And Eyepieces) and the light bulb icon (Sample Illumination), respectively

- Roughly focus on the sample by rotating the Z position dial

- Choose the objective lens (x20, x50, x100 whatever you want) by rotating the revolver

- Accurately focus on the sample by using the ball type mouse (you can use the laser spot (laser power<5x10-5))

- Click the “Measurement  New  Spectral acquisition”

- Select or enter the Grating scan type: Static or Extended, Confocality: Standard (default) or Confocal , Spectrum rage, laser (532 nm) and grating (1800 l/mm)

- Go to the “Acquisition” tab

- Set the exposure time, accumulation, and laser power

- Check the box for the “Close laser shutter on completion”

- Click “OK”

- Close the curtain and turn off the room light

- Click the “Run icon”

- File  Save file as  name your data  save as wxd. file

- Take out your sample

(Mapping)

- Click the “Measurement  New  Map image acquisition”

- Choose the mapping area by dragging in the video view (if you want high speed mapping, check off the box of “Stream Spot”)

- Choose the step size

- Set the measurement properties as described in measurement part

(Streamline mapping: high speed mapping)

- Click the “Measurement  New  Streamline image acquisition”

- Choose the mapping area (vertically long rectangular) by dragging in the video view

- Set the measurement properties as described in measurement part

Shutdown

- Close the Wire 3.2 window

- Turn off the laser (turn off the key, and then power)

- Turn off the main power

- Put the plastic cover on the Raman microscope

- Turn off the PC display